Thermo Shaker—TS100,TSC10

Description:

TS100/TCS10 Thermo Shaker is a temperature controlled shaker, which is useful for a variety of molecular biology applications. Speed, time and temperature settings are continuously visible on the LCD, simultaneously showing both actual and selected values. Integral over-temperature control ensures long time, safety and sample integrity.

Features:

- LCD display. It is easy to setup and use.

- Reliable design with inner extra temperature protection.

- Each device is equipped with lifter that is easy to replace different blocks.

- Digital accuracy and convenience.

- Precision wells for uniform thermal transfer.

- Assorted blocks for a wide variety of tube sizes.

- Brushless DC Motor: noiseless long service life maintenance-free.

- Conforms to CE safety standard.

- Accurately control and display time, temperature and speed.

- 12-month warranty.

Applications:

- Maintaining 14°C for ligation reactions.(TCS10)

- Maintaining 17°C for storing oocytes.(TCS10)

- Storing samples at ice bucket temperatures.

- Storing enzymes or DNA libraries at your work station.

- Transporting samples from the field at exactly the right temperature.

- Heating samples in the field.

- Enzyme reactions and deactivations.

- PCR sample preparation.

Parameters:

| Model | TS100 | TCS10 |
| --- | --- | --- |
| Temp setting range | 0℃~100℃ | 0℃~100℃ |
| Temp control range | RT+5℃~100℃ | RT.-20℃~100℃ |
| Time range | 1min ~ 99h59min/∞ | 1min ~ 99h59min/∞ |
| Temp accuracy | ≤±0.3℃* | ≤±0.3℃* |
| Display accuracy | 0.1℃ | 0.1℃ |
| Temp uniformity | ≤±0.3℃* | ≤±0.3℃* |
| Speed range | 200~1800rpm | 200~1500rpm |
| Orbit | 3mm | 3mm |
| Heating time | ≤12min(from 25℃ to 100℃)* | ≤15min(from 25℃ to 100℃)* |
| Cooling time1 |  | ≤10min(from 100℃ to 25℃)* |
| Cooling time2 |  | ≤20min(from RT. to RT.-20℃)* |
| Multi point running | Yes(Max 5points) | Yes(Max 5points) |
| Multi cycle | Yes(Max 99times) ） | Yes(Max 99times) |
| Auto heating | Yes | Yes |
| Auto running | Yes | Yes |
| Auto resume to run | Yes | Yes |
| Storage at 4℃ |  | Yes |
| Power supply | 150W | 150W |
| Voltage | AC220V or AC110V/50-60Hz | AC110V~220V/50-60Hz |
| Fuse | 250V 3A Ф5×20 | 250V 3A Ф5×20 |
| Dimension(mm) | 260x195x150 | 260x195x150 |
| Weight | 7.5kg | 7.0kg |

| Type | Capacity | Temp range | Max Speed (TS100) | Max Speed (TCS10) |
| --- | --- | --- | --- | --- |
| MD30 | 96x0.2ml | 0℃~100℃ | 1800rpm | 1500rpm |
| MD31 | 54x0.5ml | 0℃~100℃ | 1800rpm | 1500rpm |
| MD32 | 35x1.5ml | 0℃~100℃ | 1800rpm | 1500rpm |
| MD33 | 35x2.0ml | 0℃~100℃ | 1800rpm | 1500rpm |
| MD34 | 20x1.5ml+15x0.5ml | 0℃~100℃ | 1800rpm | 1500rpm |
| MD35 | 96x0.2ml | 0℃~100℃ | 1800rpm | 1500rpm |
| MD36 | 24x5ml | 0℃~100℃ | 1000rpm | 1000rpm |
| MD37 | 15x10ml | 0℃~100℃ | 1000rpm | 1000rpm |
| MD38 | 12x15ml | 0℃~100℃ | 1000rpm | 800rpm |
| MD39 | 6x50ml | 0℃~100℃ | 800rpm | 800rpm |

Optional Blocks: