HANGZHOU RUICHENG INSTRUMENT CO., LTD.

Hybridization Oven—HD1000
[image: image2.png]ruICIHENG




Description
The Hybridization Oven is equipped with step motor and PID intelligent temperature control technology. Its PID fuzzy control technology can ensure temperature control accuracy and automatically adjust the heating rate to save waiting time. It is widely used in the screening of cloned genes, the production of enzyme cutting maps, the qualitative and quantitative detection of specific gene sequences in genome, and the diagnosis of diseases. Therefore, it has a wide range of applications not only in the field of molecular biology but also in clinical diagnosis.

Features:

Colorful LCD display is more intuitive. It is easy to setup and use.

Excellent temperature accuracy and uniformity. 
Step motor: noiseless, less interference and maintenance-free.
It has a glass door for observing samples during processing. 
Interior surfaces are of stainless steel for high durability and low maintenance.
It can hold different kinds of holder and platform.
An alarm voice sounds automatically at the end of the run.

Built-in temperature calibration function. 
Built-in over-heating protection device, safe and reliable. 
Parameters:

	Model  HD1000
SH2000

	Temp. setting range
	0(C~100(C

	Temp. control range
	RT.+5(C~100(C

	Time range
	1min~99h59min/∞

	Temp. accuracy
	≤±0.3(C@37(C

	Display accuracy
	0.1(C

	Temp. uniformity(inside oven) 
	≤±0.3(C@37(C

	Temp. uniformity(inside bottles)
	≤±0.1(C@37(C

	Speed range
	5~30rpm

	Max. load
	5kg

	Program number
	99programs

	Auto preheating
	Yes

	Auto constant temp. at the end of run
	Yes

	Auto resume to run
	Yes

	Power supply
	1000W

	Voltage
	AC100-240V/50-60Hz

	Fuse
	250V 3A  Ф5×20

	Interior Dimension(mm)
	267x326x273

	Overall Dimension(mm)
	428x400x470

	Weight(kg)
	25


Holders and platform
	Name
	Model
	Description
	Remark

	Hybridization bottle holder
	HD1000-T1
	10pcs 35x300mm bottles or 20pcs 35x150mm bottles
	Standard

	Centrifuge tube holder
	HD1000-T2
	22tubes of 1.5ml/2.0ml(one plate)

8tubes of 10ml/15ml(one plate)
6tubes of 50ml(one plate)
	Optional, holder holds 4 plates

	Rocking platform
	HD1000-T3
	Platform dimension: 270x215mm
	Optional


[image: image1.jpg]=

2 ,
[l





Hangzhou Ruicheng Instrument Co., Ltd.

Add：Building E, No.8 of Houshan Road, Yuhang District, Hangzhou City, Zhejiang Province, China  

Tel：0571-88630940  Fax：0571-85126370

Email: info@ruiwell.com, sales2@ruiwell.com  Website: http://www.ruiwell.net

