HANGZHOU RUICHENG INSTRUMENT CO., LTD.

Dry Bath Incubator

DH2000 series
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Dry bath incubator is a sample incubation equipment based on film heating and PID automatic temperature control technology. Flexible timing, constant temperature function and unique multi-point operating mode improve the efficiency of the laboratory work, which is the ideal equipment for samples incubation, catalytic reaction process, saving and etc.

Features:

Colorful LCD display(320x480 Pixel), easy to set up and use. 
Accurate temperature control, high stability.

Different blocks for choice. Customized blocks are accepted.
Temperature calibration function.

Three kinds of countdown trigger modes available.

External temperature sensor can be connected.
Magnetic dust-proof lid can be optionally supplied(for DH2000-1 or DH2000-2).
Built-in software and hardware over temperature protection device, will use more reliable.
Parameters:

	Model
	DH2000-1
	DH2000-2
	DH2000-4

	Temp. setting range
	0℃~160℃
	0℃~160℃
	0℃~130℃

	Temp. control range
	RT+5℃~160℃
	RT+5℃~160℃
	RT+5℃~130℃

	Time range
	1min~99h59min/∞
	1min~99h59min/∞
	1min~99h59min/∞

	Temp. accuracy
	≤±0.1℃@37℃

≤±1℃@120℃
	≤±0.1℃@37℃

≤±1℃@120℃
	≤±0.2℃@37℃

≤±1℃@120℃

	Display accuracy
	0.1℃
	0.1℃
	0.1℃

	Temp. uniformity
	≤±0.1℃@37℃
	≤±0.1℃@37℃
	≤±0.2℃@37℃

	Heating rate
	1(C/min, 3(C/min, 5(C/min, AUTO
	1(C/min, 3(C/min, 5(C/min, AUTO
	1(C/min, 3(C/min, 5(C/min, AUTO

	Heating time(AUTO)
	≤25min(from 25℃ to 160℃)
	≤25min(from 25℃ to 160℃)
	≤20min(from 25℃ to 130℃)

	Program number
	99programs
	99programs
	99programs

	Multi points running
	Yes(Max 5 points)
	Yes(Max 5 points)
	Yes(Max 5 points)

	Multi cycle
	Yes(Max 99 times)
	Yes(Max 99 times)
	Yes(Max 99 times)

	Auto preheating
	Yes
	Yes
	Yes

	Auto running
	Yes
	Yes
	Yes

	Auto resume to run
	Yes
	Yes
	Yes

	Auto constant temp. at the end of run 
	Yes
	Yes
	Yes

	External temperature sensor interface 
	1port interface
	1port interface
	1port interface

	Block quantity
	1
	2
	4

	Power
	200W
	400W
	500W

	Voltage
	AC220V or AC110V/50-60Hz
	AC220V or AC110V/50-60Hz
	AC220V or AC110V/50-60Hz

	Fuse
	250V  3A  Ф 5×20
	250V  3A  Ф 5×20
	250V  3A  Ф 5×20

	Dimension(mm)
	320x205x110
	320x205x110
	385x205x110

	Weight(kg)
	4.8
	6.1
	8.8


Optional Blocks:

	Model
	Tube Diameter
	Capacity
	Model
	Tube Diameter
	Capacity

	MD01
	6mm
	48
	MD10
	26mm
	8

	MD02
	7mm
	48
	MD11
	28mm
	4

	MD03
	10mm
	24
	MD12
	40mm
	3

	MD04
	12mm
	24
	MD13
	0.2ml centrifuge tube
	48

	MD05
	13mm
	24
	MD14
	0.5ml centrifuge tube
	48

	MD06
	15mm
	16
	MD15
	1.5ml centrifuge tube
	24

	MD07
	16mm
	16
	MD16
	2.0ml centrifuge tube
	24

	MD08
	19mm
	12
	MD17*
	0.2ml PCR plate
	96

	MD09
	20mm
	12
	MD18*
	0.2ml Elisa plate
	Flat bottom


*MD17 and MD18 is special for DH2000-2/DH2000-4
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Hangzhou Ruicheng Instrument Co., Ltd.

Add：Building E, No.8 of Houshan Road, Yuhang District, Hangzhou City, Zhejiang Province, China  

Tel：0571-88630940  Fax：0571-85126370

Email: info@ruiwell.com, sales2@ruiwell.com  Website: http://www.ruiwell.net

